1.0 	Objective:

To describe the written standard procedure for Calibration & Operation of HPLC LC- 2050C

2.0 	Scope:

This SOP is applicable for Calibration & Operation of HPLC LC-2050C

# 3.0 	REFERENCE

In-House, Instrument Manual

4.0 	Responsibility:

4.1 	Officers and Executive – F&D

4.1.1 	Proper Operation of HPLC, LC- 2050C

4.1.2 	Calibration as per frequency & method.

4.1.3 	To give training as per written approved procedure by time to time.

4.2 	Engineering Department

4.2.1 	Maintenance of HPLC as per intimation from Analytical Development Department.

5.0 	Accountability:

5.1 	Manager- AD

6.0 	Procedure :

6.1 	Instrument Details:

HPLC Auto sampler - Instrument ID-XXXX & AD/LC/XXX  Model-LC-2050C & LC2050

Make-Shimadzu.

6.2 	General About Instrument :-

6.2.1 	Gradients pump.

6.2.2 	Thermostat column compartment.

6.2.3 	Auto sampler & Injector.

6.2.4 	Detector with wave length.

6.2.5 	Computer with window 10 (Pro) bases for LC-2050C with Lab Solution Software.

6.2.6 	Solvents reservoir Bottle A, B, C, D. Rinse and seal wash.

6.2.7 	 Store the instrument clean. Always keep the instrument in AC room.

6.3 	Basic Start UP Procedure For HPLC Systems:-

6.3.1 	Ensure that the system is connected to stabilize with UPS Power supply.

6.3.2 	Switch ON the main for power supply.

6.3.3 	Switch ON the main for instruments screen show.

6.3.4 	wait for Ok & asking  user password ,  	 	6.3.5 	Switch ON the pump & oven if require.

6.3.6 	Switch ON Detector.

6.3.7 	Computer ON opens with Lab solution Software.

6.4 	Lab Solution Software Operation:

6.4.1 	Cursor put on Lab solution icon & Double Click.

6.4.2 	Enter the user ID and password, Click on ‘OK’

6.4.3 	Double click on the instrument ID which is required.

6.4.4 	Now Lab solution software open for operation & come beep sound 2 times.

6.4.5 	Open new method file & generate new method for new product

6.4.6 	Close method file last open, before open new method file (existing)  	 	6.4.7 	Post rum File: for running, existing data graphs observation & printing.

6.5 	Data Acquisition Parameters:-

6.5.1 	Data Acquisition: LC time process (sample run time)  	 	6.5.2 	Gradient time Process: Sample gradient analysis time setting.

6.5.3 	Flow Rate: 0.2 to 2.0 ml, & Down Load.

6.5.4 	M.P. Lines A. B, C, D tubes Volume set in %.

6.5.5 	Maximum Pressure set (as require Max: 440)

6.5.6 	Set Lamp ON/OFF & Down Load

- Temp: L/H.

- Wave Length selection Like 254 nm.

- Volt   : OK

- Range: 1.0 %

6.5.7 	Oven: Column Temperature selection 200C to 700C as require & in degases.

6.5.8 	 Auto Sampler:-

6.5.8.1 	Rack 1, 2, 3 or 4  : Select vials 1.5 ml x 54, when D2 on

6.5.8.2 	Rinsing            : 200 µl

6.5.8.3 	Needle             : 52 mm  	 	 	6.5.8.4 	SMP Speed     : 15 ul / sec.

6.5.8.5 	Rinse Needle.  : Before & after

6.5.9 	Auto Purging:-

6.5.9.1 	Purging Order: A, B, C, D and rinse port.

6.5.9.2 	Mobile Phase  : A, B,C, D.

6.5.9.3 	Purging Time   : 1….to….10 min.

6.5.9.4 	Purging Pump : LC-2030C & Plus

6.5.10 (Window open for new method & new product file (Unknown file Single & Multiple

Injection):-

6.5.10.1 	For Window: - Change in wizard

6.5.10.2 	New Batch file Sequences

6.5.10.2.1 	Show window -   Batch file Wizard  Method show   (XYZ)

Tray No injection volume  No of samples   (New) unknown only   sample name sample ID data file  finish.

6.6 	Calibration Procedure of HPLC:-

6.6.1 	A] Calibration of Pump:-

6.6.1.1 	Calibration of Pump is to be done with the help of distilled water at 250C±20C.

6.6.1.2 	Set the flow of the pump at 1ml/min. keep the system at this flow rate for about 10 minutes to equilibrate the system without the Column.

6.6.1.3 	Now dry to A grade measuring cylinder & weight the empty A grade measuring cylinder

before starting the procedure.

6.6.1.4 	Start the stopwatch & simultaneously start collecting the eluent from the pump outlet in a class A grade measuring cylinder (10ml capacity).

6.6.1.5 	Measure the time it takes for water to fill up to the graduation marked and calculate the flow accuracy and flow rate.

6.6.1.6 	Calculate the  Flow rate  accuracy by following formula

Flow rate accuracy (%) = (actual Flow rate/Set Flow rate-1)x100

Set time 0.5 ml/min 1200(s), 1.0 ml/min 600(s), 2.0 ml/min 300(s),

6.6.1.7 	Calculate the Actual Flow rate by following formula

Actual Flow rate = Set flow rate X set time (s)/actual measurement time

Set time 0.5 ml/min 1200(s), 1.0 ml/min 600(s), 2.0 ml/min 300(s),

6.6.1.8 	Set the flow 0.5, 1.0, & 2.0 ml & repeat the above experiment, allowing the system to

equilibrate at least for 10 minutes at each flow set.

6.6.1.9 	Acceptance Criteria:

For: 0.5 ml/min  ± 0.010 ml/min [0.490-0.510 ml/min]

For: 1.0 ml/min  ± 0.020 ml/min [0.980-1.020 ml/min]

For: 2.0 ml/min  ± 0.040 ml/min [1.960-2.040 ml/min]

6.6.1.10 	Same way Calibrate A, B, C, D pump individually as above procedure.

6.6.1.11 	B) Noise & Drift Calibration For UV;(Limit : Noise: 20 x10-3 AU/h & Drift : 500 x10-3 AU/h

For PDA:  Limit : Noise: 50 x10-3 AU/h & Drift : 500 x10-3 AU/h  	 	 	 	Pre-treatment procedure: for System Before Calibration.

- Sonicate suction filter for 10 minutes in 10% v/v Nitric acid and then 5 minutes in 100 % methanol.

- Set the each 4 port with 25% ratio and run HPLC grade IPA 100% for 30 min with flow rate 5ml / min.

- Deep all ports in HPLC grade Hot water & purge all lines.

- Connect the tubes with union & Replace the Column.

- Set the each 4 port with 25% ratio and run HPLC grade hot water for 30 min with flow rate 5ml /min.

Condition  for drift and Noise:

Mobile phase     : HPLC grade water (Dip all port in water) Flow rate           : 1.0 ml/min.

Wave Length    : 250 nm,

Column Oven   : 25°C

Column: connect with union or capillary column.

Saturation time:  60 minutes

Run Time:  15 minutes

6.6.1.12 	C) Gradient Pump Calibration (Limit + 2.0%)

| Mobile Phase | A Line & C Line | Water |
| --- | --- | --- |
|  | B Line & D Line | 10 ppm Caffeine in water |
| Chromatographic condition | Flow | 2 ml/min |
|  | Wavelength | 272 nm |
|  | Run time | 26 Min |
|  | Vial | System Blank |

| Min | A | B | C | D |  |
| --- | --- | --- | --- | --- | --- |
| 1 | 50 | 0 | 50 | 0 |  |
| 1.01 | 40 | 10 | 40 | 10 |  |
| 4 | 40 | 10 | 40 | 10 |  |
| 4.01 | 30 | 20 | 30 | 20 |  |
| 7 | 30 | 20 | 30 | 20 |  |
| 7.01 | 20 | 30 | 20 | 30 |  |
| 10 | 20 | 30 | 20 | 30 |  |
| 10.01 | 10 | 40 | 10 | 40 |  |
| 13 | 10 | 40 | 10 | 40 |  |
| 13.01 | 0 | 50 | 0 | 50 |  |
| 16 | 0 | 50 | 0 | 50 |  |
| 21 | 50 | 0 | 50 | 0 |  |
| 26 | 50 | 0 | 50 | 0 | STOP |

6.6.1.12 	Determine Lamp Energy: NLT 400 mv by Lab solution software.

6.6.2 	Vial Position Test: Use linearity solution to perform this test, for different position of sample injection in tray as below and calculate the linearity.

| V | ial Position at tray |  |
| --- | --- | --- |
| 1-1 | 3-33 | 3-54 |
| 2-9 |  | 4-46 |

6.6.3 	Calibration of Detector:-

6.6.3.1 	Repeatability (Precision)

6.6.3.2 	Absorbance linearity

6.6.3.3 	System Repeatability of Peak area and Retention time:-

System Parameters    :

Column             	   : C18  250 X 4.6 mm, 5 µ

Mobile phase                     : Water: Methanol = 60:40

Flow rate 	 	  : 1.00ml/min

Wavelength   	 	  : 272 nm

Column oven Temp         : 40°C

Volume  	 : 10 µl

Analysis Time                  : 3-8Min

Caffeine Solution            : 20mcg/ml in MP: Separately 6 replicates Injections.

6.6.3.4 	Linearity of Detector absorbance:

System Parameters:  same as above system & condition:

Caffeine solution     : 10/15/20/25/30 mcg/ml in MP.

Caffeine anhydrous Stock solution: Caffeine 20mg--------100ml

| Sr. No. | Diluted Solution | Conc. mcg/ml | Area | R square NLT 0.99 |
| --- | --- | --- | --- | --- |
| 01 | 1ml---20 | 10 |  |  |
| 02 | 1.5 ml---20 | 15 |  |  |
| 03 | 2 ml---20 | 20 |  |  |
| 04 | 2.5ml---20 | 25 |  |  |
| 05 | 3ml---20 | 30 |  |  |

6.6.4 	Calibration of Injector (Loop):-

6.6.4.1 	Chromatographic Condition:-

Column                : C18  250 X 4.6 mm, 5 µ

Mobile phase       : Methanol : Water (40 : 60)

Flow Rate            : 1.0 ml/min

Wavelength         : 272nm

6.6.4.2 	Preparation of Standard Solution:-

6.5.4.2.1 Weigh accurately 20 mg of caffeine into a 100ml volumetric flask, dissolve

& dilute to volume with water. Further dilute 10 ml. of this solution to 100 ml with water. (20 mcg/ml.)

6.5.4.2.2 Inject 10µl, 20µl, 50µl of the standard solutions five replicates & record the chromatograms.

6.5.4.2.3 Calculate the %RSD of the peak area of caffeine from the 5 replicate injections at each level.

6.5.4.2.4 Calculate the average peak area of caffeine from the 5 replicate injections at

each level.

6.5.4.2.5 Acceptance Criteria:-

% RSD limit for 5 replicate injection samples should NMT 2.0%.

6.6.5 	Calibration of Thermostat for column oven:-

6.6.5.1 	Purge the system with purified water to remove any solvents & previous buffer salts.

6.6.5.2 	Keep the thermometer in the right side of the column oven.

6.6.5.3 Set the temperature at 20°C. Wait till the set temperature is attained & the temperature display on the instrument is stable.

6.6.5.4 	Record the temperature displayed by the instrument & the thermometer.

6.6.5.5 	Repeat the steps for temperature setting at 30°C, 40°C, 50°C, 70°C.

6.6.5.6 	Record the observation.

6.6.5.7 	Record the temperature displayed by the instrument & the thermometer.

6.6.5.8 Acceptance Criteria: - The difference between the set temperature & the displayed temperature on the thermometer should not be more than ±1°C.

6.6.6 	Wavelength Accuracy Calibration :- (Auto sampler HPLC)

6.6.6.1 	Verify the wavelength accuracy use 20 mcg/ml dilution of caffeine in water.

Caffeine solution-   ……. mg             100 ml         10 ml         50 ml

6.6.6.2  Mobile Phase - Methanol: Water (40: 60).

Column           - C18  250 X 4.6 mm, 5 µ

6.6.6.3  Flow Rate        - 1 ml/min

Primary Std. No. ………….

6.6.6.4  Inject the solution & enter the value of wave length 269 nm to 275 nm.

Note down the areas observed. And other wavelength 202 nm to 208nm

Note down the areas observed. And other wavelength 241 nm to 247 nm

6.6.6.5 	For PDA detector used 1st injection of Repeatability test and scan with spectra.

6.6.6.6  Acceptance Criteria - Maxima 273± 2 nm

Acceptance Criteria - Minima 243 ± 2 nm

Acceptance Criteria - Maxima 204± 2 nm

6.6.7 	Carry over Test (NMT 0.01%)

6.6.7.1 For carry over test first inject water as blank followed by three injections of 200 mcg/ml dilution of caffeine in mobile phase & then blank inj.

Calculation:

% Carry Over = (Area of caffeine in blank 2 - Area of caffeine in blank 1)x100                                  Mean area of caffeine in sample solution

6.6.8 	Frequency of Calibration: -

6.6.8.1 Calibration shall be carried out in every six months or if any major break down observed in specific parts (hardware),the calibration shall be carried for same related part before use the same instrument as below mention table :-

| Sr. No | Name of HPLC Parts under breakdown | Related test for calibration after breakdown |
| --- | --- | --- |
| 1 | Pump | Pump calibration |
| 2 | D2 Lamp | Wavelength accuracy |
| 3 | Solenoid valve | Calibration of GPV |
| 4 | Detector | Detector calibration point & test Repeatability test |
| 5 | Injector Loop | Injection Volume Precision (Calibration of injector) |

6.6.8.2 	Breakdown shall be reported through laboratory incident ref. to Respective SOP.

6.6.8.3 After calibration record all the chromatograms and results filled to Annexure I. If all result are within the SOP specification limit than use the same HPLC system for routine analysis.

6.6.9 	Calibration of Sample Cooler:-

6.6.9.1 	Purge the system with purified water to remove any solvents & previous buffer salts.

6.6.9.2 	Keep the thermometer in the right side of the column oven.

6.6.9.3 	Set the temperature at 5°C. Wait till the set temperature is attained & the temperature

display on the instrument is stable.

6.6.9.4 	Record the temperature displayed by the instrument & the thermometer.

6.6.9.5 	Repeat the steps for temperature setting at 10°C, 15°C, 20°C, 25°C.

6.6.9.6 	Record the observation.

6.6.9.7 	Record the temperature displayed by the instrument & the thermometer.

6.6.9.8 Acceptance Criteria: - The difference between the set temperature & the displayed temperature on the thermometer should not be more than ±1°C.

7.0 	Abbreviations:

7.1 	SOP 	: Standard Operating Procedure

7.2 	QA 	: 	 Quality Assurance

7.3 	HPLC : High Performance Liquid Chromatography

8.0 	Annexure (s):                                                                             Format No.:

Annexure I – Calibration of HPLC Auto Sampler                    F/AD/044/A/00  	 	Annexure II –HPLC User Log Book                                       F/AD/044/B/00

SUMMARY OF HPLC CALIBRATION

Annexure - I 	 	 	 	 	     	             Format No.: F/AD/044/A/00

| Instrument Name: | HPLC AUTO SAMPLER | Model Number: | LC2050C |
| --- | --- | --- | --- |
| Serial number: |  | Instrument ID number: | AD/LC/ |
| Date of Calibration |  | Next Calibration due date: |  |

| Sr. No. | Test Performed |  | Acceptance criteria | Observation |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | A] Flow accuracy of pump (ml/min) |  | Tolerance (mL/min) | A | B | C |  | D |
|  | 0.5 |  | 0.490 to 0.510 |  |  |  |  |  |
|  | 1.0 |  | 0.980 to 1.020 |  |  |  |  |  |
|  | 2.0 |  | 1.960 to 2.040 |  |  |  |  |  |
|  | B) Calibration of Noise & Drift | For UV | Noise 20 x10-3 AU/h Drift 500 x10-3 AU/h | Noise : Drift : |  |  |  |  |
|  |  | For PDA | Noise 50 x10-3 AU/h) Drift 500 x10-3 AU/h | Noise : Drift: |  |  |  |  |
|  | C] Calibration of GPV |  |  |  |  |  |  |  |
|  | 0.0 % |  | NA |  |  |  |  |  |
|  | 20 % |  | 18 % - 22 % |  |  |  |  |  |
|  | 40 % |  | 38 % - 42 % |  |  |  |  |  |
|  | 60 % |  | 58 % - 62 % |  |  |  |  |  |
|  | 80 % |  | 78 % - 82 % |  |  |  |  |  |
|  | 100 % |  | 98 % - 102 % |  |  |  |  |  |
|  | UV Lamp Energy |  | NLT 400 mv |  |  |  |  |  |
| 2 | DETECTOR CALIBRATION TEST |  | Area -NMT 2 .0 % |  |  |  |  |  |
|  | A) Repeatability test. |  | R-time-NMT 1.0 % |  |  |  |  |  |
|  | B) Absorbance Linearity Test. |  | R2 – NLT: 0.997 % |  |  |  |  |  |
|  | C) Vial Position Test. |  | Complies |  |  |  |  |  |
| 3 | Injection Volume Precision (Calibration of injector) |  | % RSD= NMT 2.0 | Inj.Vol. | % RSD-Area |  | % RSD-RT |  |
|  |  |  |  | 10 µl |  |  |  |  |
|  |  |  |  | 20 µl |  |  |  |  |
|  |  |  |  | 50 µl |  |  |  |  |
| 4 | Thermostat Oven |  | NMT ± 1° C |  |  |  |  |  |
| 5 | Carry over Test |  | 0.01 % |  |  |  |  |  |
| 6 | Wavelength Accuracy |  | Maxima 273 ± 2 nm |  |  |  |  |  |
|  |  |  | Maxima 204 ± 2 nm |  |  |  |  |  |
|  |  |  | Minima 243 ± 2 nm |  |  |  |  |  |
| 7 | Sample Oven |  | NMT ± 1° C |  |  |  |  |  |

CALIBRATION STATUS: The calibration of the instrument is within limit/ out of limit.

Calibrated By Sign/Date:                                                                            Checked By Sign/Date:

CALIBRATION OF PUMP: (Dist water Temp at 250C - 270C)

Set the flow 0.5, 1.0 & 2.0 ml & repeat the above experiment, allowing the system to equilibrate at least for 10 minutes at each flow set.

STOP WATCH ID No: ……………………………. Calibration due date: ………………………..

Acceptance criteria:  ± 2.0 % in ml

Pump. A

| Set Flow rate (ml/min) | Set Time (s) | Actual Measurement time | Flow Rate Accuracy | Actual Flow rate |
| --- | --- | --- | --- | --- |
| 0.5 ml | 1200 |  |  |  |
| 1.0 ml | 600 |  |  |  |
| 2.0 ml | 300 |  |  |  |

Pump. B

| Set Flow rate | Set Time (s) | Actual Measurement time | Flow Rate Accuracy | Actual Flow rate |
| --- | --- | --- | --- | --- |
| 0.5 ml | 1200 |  |  |  |
| 1.0 ml | 600 |  |  |  |
| 2.0 ml | 300 |  |  |  |

Pump. C

| Set Flow rate | Set Time (s) | Actual Measurement time | Flow Rate Accuracy | Actual Flow rate |
| --- | --- | --- | --- | --- |
| 0.5 ml | 1200 |  |  |  |
| 1.0 ml | 600 |  |  |  |
| 2.0 ml | 300 |  |  |  |

Pump. D

| Set Flow rate | Set Time (s) | Actual Measurement time | Flow Rate Accuracy | Actual Flow rate |
| --- | --- | --- | --- | --- |
| 0.5 ml | 1200 |  |  |  |
| 1.0 ml | 600 |  |  |  |
| 2.0 ml | 300 |  |  |  |

Calibrated by (Sign/Date):                                                                     Checked by (Sign/Date):

Calibration of Noise and Drift:

| Detector | Noise Limit | Drift: | Observed Result |  |
| --- | --- | --- | --- | --- |
|  |  |  | Noise | Drift |
| PDA | Noise 50 x10-3 AU/h | 500 x10-3 AU/h |  |  |
| UV | Noise 20 x10-3 AU/h | 500 x10-3 AU/h |  |  |

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

Gradient Pump Calibration (Limit +/- 2.0%)

| Quaternary gradient system | Limit | Observation (result) |
| --- | --- | --- |
| 0 % | NA |  |
| 20 % | 18 % - 22 % |  |
| 40 % | 38 % - 42 % |  |
| 60 % | 58 % - 62 % |  |
| 80 % | 78 % - 82 % |  |
| 100 % | 98 % - 102 % |  |

Results:  complies / not complies

UV Lamp Energy test: NLT 400 MV :…………………………..

Calibrated by (Sign/Date):                                                                        Checked by (Sign/Date):

2] DETECTOR CALIBRATION POINT & TEST

(a) REPEATABILITY TEST             (b) ABSORBANCE LINEARITY TEST

a] System repeatability of peak area and retention time (Limit -RSD NMT 1.0 %)

System Parameters           : 1st wash with Methanol.

Column ID                       :…………………………….. (C18 250 X 4.6 mm, 5µ)

Mobile phase                       : Water: Methanol = 60:40

Flow rate 	 	     : 1.00ml/min

Wavelength   	     	     : 272 nm

Column oven Temp            : 40°C

Volume                               : 10 µl

Analysis time                      : 3-8Min

Caffeine solution                : 20mcg/ml in MP: Separately 6 replicates injections.

Test Preparation: (Caffeine ……… mg-------ml-…………………….)

Primary Std Batch No - ………………………    Opened date:  ……………  Exp. Date:   ………

OR W.S No._____________  	 	 	   Vial No.____________

| Sr. No. | Conc. | Area | %RSD | Retention Time | %RSD |
| --- | --- | --- | --- | --- | --- |
| 01 | 20 mcg/ ml |  |  |  |  |
| 02 | 20 mcg/ ml |  |  |  |  |
| 03 | 20 mcg/ ml |  |  |  |  |
| 04 | 20 mcg/ ml |  |  |  |  |
| 05 | 20 mcg/ ml |  |  |  |  |
| 06 | 20 mcg/ ml |  |  |  |  |
|  | Mean |  |  |  |  |

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

b] LINEARITY OF DETECTOR RESPONSE (Absorbance linearity Test)

System Parameters:  Same as above system & condition:

Caffeine solution         : 10/15/20/25/30 mcg/ml in water

Caffeine Stock Solution: (Caffeine …………mg--------……..ml)

Primary Std Batch No - ………………………    Opened date:  ……………  Exp. Date:   ………

OR W.S No._____________  	 	 	   Vial No.:____________

| Sr. No. | Diluted Solution | Conc. mcg/ml | Area | R square NLT 0.997 |
| --- | --- | --- | --- | --- |
| 01 | 1ml---20 | 10 |  |  |
| 02 | 1.5ml---20 | 15 |  |  |
| 03 | 2ml---20 | 20 |  |  |
| 04 | 2.5ml---20 | 25 |  |  |
| 05 | 3ml---20 | 30 |  |  |

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

3] CALIBRATION OF INJECTOR (loop):

Chromatographic condition:

Column ID     - …………………………….. (C18 250 X 4.6 mm, 5µ)

Mobile phase       	: Water: Methanol = 60:40

Flow rate 	 	: 1.00 ml/min

Wavelength   	  	: 272 nm

Column oven Temp      : 40°C

Caffeine standard solution: 20 mcg/ml in water

Primary Std:  Batch No - ……………………… Opened date:  ……………  Exp. Date:   ………

OR W.S No._____________  	Vial No.____________

Inject 10µl, 20µl, 50µl, of the standard solutions five replicates & record the chromatograms.

Calculate the %RSD of the peak area of caffeine from the 5 replicate injections at each level.

Calculate the average peak area of caffeine from the 5 replicate injections at each level.

Acceptance criteria :( % RSD: NMT 2.0)

| Sr. No. | 10µl area | 20µl area | 50µl area | % RSD |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Inj. Volume | Area | RT |
| 01 |  |  |  | 10 µl |  |  |
| 02 |  |  |  |  |  |  |
| 03 |  |  |  | 20 µl |  |  |
| 04 |  |  |  |  |  |  |
| 05 |  |  |  | 50 µl |  |  |
| Average |  |  |  |  |  |  |

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

- CALIBRATION OF THERMOSTAT COLUMN OVEN:

Thermometer Id No: …………………………Range:   …………..   Calibration Due date:………

Set the temperature at 20°C. Wait till the set temperature is attained & the temperature display on the instrument is stable.. Record the temperature displayed by the instrument & the thermometer. Repeat the steps for temperature setting at 20°C, 30°C, 40°C, 50°C & 70°C.Record the observation.

Acceptance criteria: (NMT +/- 1° C)

| Sr. No. | 20°C | 30°C | 40°C | 50°C | 70°C |
| --- | --- | --- | --- | --- | --- |
| 01 |  |  |  |  |  |
| 02 |  |  |  |  |  |
| Mean |  |  |  |  |  |

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

- CARRY OVER TEST:

Primary Std:  Batch No - ……………………… Opened date:  ……………  Exp. Date:   ………

OR W.S No._____________  	Vial No.____________

Sample Preparation: - Prepare 200 mcg/ml of caffeine in water.

Caffeine solution-   ……. mg   ---->…….. ml ----->………ml  ----> ……… ml

| Sr. No. | Mean Area of caffeine in sample solution | Area of in blank | caffeine injection | % Carry over | Limit |
| --- | --- | --- | --- | --- | --- |
|  |  | Blank 1 |  |  | 0.01% |
|  |  | Blank 2 |  |  |  |

Calculation: % Carry Over = (Area of caffeine in blank 2 - Area of caffeine in blank 1) x100

Mean area of caffeine in sample solution

Calculation:

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

- WAVELENGTH ACCURACY CALIBRATION: (For UV detector)

To verify the wavelength accuracy use 20 mcg/ml dilution of caffeine in water. Caffeine solution-   ……. mg 100 ml           l

Mobile Phase - 40 Methanol: 60 water

Column ID     - …………………………….. (C18 250 X 4.6 mm, 5µ)

Flow rate        - 1 ml/min

Primary Std:  Batch No - ……………………… Opened date:  ……………  Exp. Date:   ………

OR W.S No._____________  	Vial No.____________

Inject the solution & enter the value of wave length 269 nm-275 nm.

Note down the areas observed.

1st Maxima       Acceptance Criteria - Maxima 273± 2 nm

| Sr. No. | Wave length | Area | Maxima |
| --- | --- | --- | --- |
| 01 | 269 nm |  |  |
| 02 | 270 nm |  |  |
| 03 | 271 nm |  |  |
| 04 | 272 nm |  |  |
| 05 | 273 nm |  |  |
| 06 | 274 nm |  |  |
| 07 | 275 nm |  |  |

2nd Maxima       Acceptance Criteria - Maxima 204± 2 nm

Inject the solution & enter the value of wave length 202 nm-208 nm. Note down the areas observed

| Sr. No. | Wave length | Area | Maxima |
| --- | --- | --- | --- |
| 01 | 202 nm |  |  |
| 02 | 203 nm |  |  |
| 03 | 204 nm |  |  |
| 04 | 205 nm |  |  |
| 05 | 206 nm |  |  |
| 06 | 207 nm |  |  |
| 07 | 208 nm |  |  |

3rd Minima       Acceptance Criteria - Minima 243 ± 2 nm

Inject the solution & enter the value of wave length 241 nm-248 nm. Note down the areas observed

| Sr. No. | Wave length | Area | Minima |
| --- | --- | --- | --- |
| 01 | 241 nm |  |  |
| 02 | 242 nm |  |  |
| 03 | 243 nm |  |  |
| 04 | 244 nm |  |  |
| 05 | 245 nm |  |  |
| 06 | 247 nm |  |  |
| 07 | 248 nm |  |  |

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

WAVELENGTH ACCURACY CALIBRATION: (For PDA detector)

Take 1st injection of repeatability and scan with spectra

| Sr. No. | Wave length | Acceptance Criteria | Result |
| --- | --- | --- | --- |
| 01 | 241 nm-248 nm | Minima 244 ± 2 nm |  |
| 02 | 202 nm-208 nm | Maxima 205± 2 nm |  |
| 03 | 269 nm-275 nm | Maxima 272± 2 nm |  |

Calibrated by (Sign/Date):                                                                Checked by (Sign/Date):

7. CALIBRATION OF SAMPLE OVEN:

Thermometer ID No.: …………………………Range:   …………..   Calibration Due date:………

Acceptance criteria: (NMT +/- 1° C)

The difference between the set temperature & the displayed temperature on the thermo-meter should not be more than ±1°C.

| Sr. No. | 5°C | 10°C | 15°C | 20°C | 25°C |
| --- | --- | --- | --- | --- | --- |
| 01 |  |  |  |  |  |
| 02 |  |  |  |  |  |
| Mean |  |  |  |  |  |

Calibrated by (Sign/Date):                                                                            Checked by (Sign/Date):

Affix the Calibration Label on Instruments:

Name of Instrument   : Shimadzu HPLC, LC-2050C

Instrument ID             : AD/LC/……

Date of Calibration     : ……………………..

Due Date (After Six month/ as when required if any breakdown): ……………………..

Calibrate by :                                                                                      Checked by :

Sign/Date     :  	 	 	 	 	                   Sign/Date    :

Approved By Manager (Sign/Date):

END OF DOCUMENTS